## Slide 1
Crystalline silicon solar cell — cross-section
incident light
front contact fingers (Ag)
anti-reflection coating
n⁺ emitter
p-type base (absorber)
p⁺ back surface field
rear metallisation (Al)
z
0
W
Wavelength selects depth: blue light is absorbed within the emitter and coating, red light reaches the rear.

## Slide 2
Where the incident photons go — EQE vs IQE
front surface
incident photons Φ₀
reflected R
coating + emitter: parasitic absorption Aₑₓₜ
Tₑₓₜ = 1 − R − Aₑₓₜ
absorber: photons absorbed → electron-hole pairs generated
collected
→ photocurrent
recombined
→ lost
EQE = collected / incident
IQE = collected / entered
IQE = EQE / (1 − R) (10)
EQE = Tₑₓₜ × IQE (11)
optics × electronics
EQE counts every incident photon, so it is charged for reflection. IQE divides that out and reports only how well the cell converts the light it actually receives.

## Slide 3
Generation depth and carrier collection
front / junction
rear surface
400 nm
700 nm
1000 nm
depth z
0
W
collection efficiency ηc
high
low
Blue light generates carriers where collection is efficient but front losses are largest; red light generates them deep, where the diffusion length and rear surface decide the outcome.

## Slide 4
Al-BSF vs PERC — the rear side sets the near-IR response
p⁺ Al-BSF
dielectric +
local contacts
Al-BSF
PERC
lost at rear
reflected back
high rear recombination (large S)
weak rear reflection
passivated rear (small S)
rear reflector returns near-IR light
Both cells share the same front optics, so their EQE and IQE curves differ only above ~900 nm — the signature of a rear-side difference.

## Slide 5
Differential spectral responsivity (DSR) measurement
Xe / halogen
lamp
mono-
chromator
chopper
beam
splitter
monitor
photodiode
monochromatic,
chopped
bias lamp
(white light)
solar cell on temperature-controlled chuck
trans-impedance
amplifier
lock-in amplifier
Δjsc (AC signal)
reference cell of known SR → sets the absolute scale
s̃ (λ, Ebias) = Δjsc / ΔEλ (15)
The bias lamp sets a realistic injection level; the lock-in recovers only the small chopped signal.

## Slide 6
How to update these figures
1 · Edit, then export
Edit any slide in this deck, then export every slide as JPEG
at about 150 dpi (File → Export → JPEG → All slides).
Simpler: run this in the EQE folder — it renders every slide
and copies the JPEGs to both places for you:
./export_figures.sh
2 · Save each JPEG in BOTH folders
Same file name in each. Miss the second one and the website
keeps showing the old picture.
EQE/figures/<name>.jpg
used by the Jupyter notebook
docs/assets/<name>.jpg
used by the documentation website
Slide → file name
Slide 1
fig_cell_structure.jpg
Slide 2
fig_optical_losses.jpg
Slide 3
fig_collection_efficiency.jpg
Slide 4
fig_albsf_vs_perc.jpg
Slide 5
fig_measurement_setup.jpg
This .pptx is deliberately not tracked by git — it is your local working copy. The exported .jpg files are what the repository and the website use, so those are the ones to commit.